## Slide 1: Thesis supervisor Agnès Bonvilain
Presented by: Adrien Mialland
ORL surgeon: Ihab Atallah
Host laboratoiry : Grenoble Image Parole Signal Automatique (Gipsa-Lab)
Université Grenoble Alpes (UGA)
Design and control of a tracheal
occlusion system for an
implantable active artificial larynx

## Slide 2: Presentation Plan
Context
Clinical Research Protocol (CRP)
Statistical Study
Real time detection
Implantable Active Artificiel larynx (IAAL)
Conclusion and perspectives

## Slide 3: 1. Context - larynx
thyroid
thyroid
larynx
larynx
trachea
trachea
œsophagus
3 main functions:
Breathing – Swallowing - Phonation
E. Maried et al. 2013
epiglotte
Vocal cordes

## Slide 4: 1. Context – Total Laryngectomy
Rehabilitation methods:
Tracheostomy.
Voice restoration.
Christopher G. Tang et al. 2015.
trachée
trachéostomie
œsophage
larynx

## Slide 5: 1. Context – Swallowing Restoration
œsophagus
trachea
Larynx

## Slide 6: 1. Context – Swallowing Restoration
View from
A et B
Debry C. et al. 2014.
œsophage
trachée
Larynx
First Artificial Larynx - 2012:
Valve mechanism - passive,
Implanted on 6 patients,
Only used under medical control,
Lack of sealing

## Slide 7: 1. Context – Swallowing Restoration
œsophage
trachea
Larynx
Time (s)
Start of
Swallowing ?
Available
Time ?
Tends to be dedicated to swallowing ?
Detection
limit
Start
End
Real time detection - constraints:
Earliness
Accuracy
Preservation
Consomption
Example of a swallowing signal

## Slide 8: Presentation Plan
Context
Clinical Research Protocol (CRP)
Statistical Study
Real time detection
Implantable Active Artificiel larynx (IAAL)
Conclusion and perspectives

## Slide 9: 2. CRP – Main criteria: earliness
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time
Time (s)
Example of a swallowing signal

## Slide 10: 2. CRP – Main criteria: earliness
Pharyngeal Muscles:
Superior constrictor
Middle constrictor
Inferior constrictor
Submandibular
muscles
Upper oesophageal sphincter
trachea
œsophagus
Hyoid Bone
epiglotte
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 11: 2. CRP – Start of Swallowing
Mylohyoid
Anterior belly of the digastric
Geniohyoid
Hyoid bone
Submandibular
Hyoid bone
Mylohyoid
Anterior bely of the digastric
[D. Park et al. 2017, J. Walton et al. 2018, R. L. Drake et al. 2010]
A. Mialland et al. Toward a robust swallowing detection for an implantable active artificial larynx: A survey. In: Med. Biol. Eng. Comput. (2023).
Upper oesophageal sphincter
Submandibular
muscles
Hyoid Bone
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 12: 2. CRP – Start of Swallowing
-200
0
200
400
mV
6.0
6.5
7.0
7.5
second
Surface electromyography electrodes
Submandibular
muscles
Hyoid Bone
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time
Upper oesophageal sphincter

## Slide 13: 2. CRP – Temporal detection limit
Before total laryngectomy
[Morinière et al. 2011, M. Hasegawa et al. 2018, T. Omari et al 2018, A. Pal et al. 2003, J. Lee et al. 2008, J. M. Dudik et al. 2015]
Pharyngeal
component
Upper oesophageal sphincter component
Laryngeal opening
component
Upper oesophageal sphincter
cricoid cartilage
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time
Submandibular
muscles
Hyoid Bone
Upper oesophageal sphincter

## Slide 14: 2. CRP – Temporal detection limit
Before total laryngectomy
[Morinière et al. 2011, M. Hasegawa et al. 2018, T. Omari et al 2018, A. Pal et al. 2003, J. Lee et al. 2008, J. M. Dudik et al. 2015]
Upper oesophageal sphincter
cricoid cartilage
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time
Submandibular
muscles
Hyoid Bone
Upper oesophageal sphincter
After total laryngectomy

## Slide 15: 2. PRC – Stylohyoid and posterior digastric muscles
Activation variability of the stylohyoid and posterior digastric muscles:
Activated at the same time than the submandibular muscles
  EMG studies on animals [Rebecca Z. German et al. 2009, A. J. Thexton et al. 2007 , R. W. Doty et al. 1956].
  Imaging study on human [T. Okada et al. 2013]
  Anatomic properties study on human [W. G. Pearson et al. 2011]
Preservation
Preservation
Consumption
Consumption
Earliness
Accuracy
Accuracy
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 16: 2. PRC – Stylohyoid and posterior digastric muscles
Preservation
Preservation
Consumption
Consumption
Earliness
Accuracy
Accuracy
Posterior Digastric (PD)
Stylohyoid (SH)
Hyoid Bone
Submandibular
muscles
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 17: 2. PRC – Stylohyoid and posterior digastric muscles
Earliness
Accuracy
Preservation
Consumption
A. J. Thexton et al. 2012
Intramuscular EMG on pigs
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 18: 2. PRC – Stylohyoid and posterior digastric muscles
Earliness
Accuracy
Preservation
Consumption
A. J. Thexton et al. 2012
Stylohyoid and posterior digastric muscles:
Predominently activate for swallowing
  EMG studies on animals [Rebecca Z. German et al. 2009, C. J. Mayerl et al. 2021, A. J. Thexton et al. 2012].
  Imaging study on human[W. G. Pearson et al. 2013]
  Anatomic properties study on human [W. G. Pearson et al. 2011]
Intramuscular EMG on pigs
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 19: 2. PRC – Stylohyoid and posterior digastric muscles
Temps (s)
D. Park et al. 2017
Earliness
Accuracy
Consumption
Preservation
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 20: 2. PRC – Stylohyoid and posterior digastric muscles
Temps (s)
D. Park et al. 2017
Earliness
Accuracy
Consumption
Preservation
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 21: 2. PRC – Stylohyoid and posterior digastric muscles
Digastrique
postérieur
stylohyoïdien
Earliness
Accuracy
Consumption
Preservation
Start of
swallowing
Detection
limit
Interval of activation of
measured muscles
Signal
processing
Actuation of
the system
Moment of activation of
Measured muscles
Time

## Slide 22: 2. PRC – Innovative measurement method
A. Mialland et al. Stylohyoid and posterior digastric measurement with intramuscular EMG, submental EMG and swallowing sound. In: BIOSIGNALS (2023)
(PD)
(SH)
(SM)

## Slide 23: 2. PRC – Innovative measurement method
17 healthy participants (8 men / 9 women)
4 swallowing: saliva, water, thick (compote), solid (madeleine).
13 non-swallowing: mouth opening, lips purseing, teeth clenching, smiling, whistling, coughing, blowing, speaking, saying « iii », jaw movements, lateral head movements, head flexion/extension, mastication.
A. Mialland et al. Stylohyoid and posterior digastric measurement with intramuscular EMG, submental EMG and swallowing sound. In: BIOSIGNALS (2023)
(PD)
(SH)
(SM)

## Slide 24: 2. PRC – Innovative measurement method
stylohyoïdien
Digastrique postérieur
Sous-mandibulaires
Son de la déglutition
Swallowing acquisition exemple

## Slide 25: Presentation Plan
Context
Clinical Research Protocol (CRP)
Statistical Study
Real time detection
Implantable Active Artificiel larynx (IAAL)
Conclusion and perspectives

## Slide 26: 3. Statistical study– Data extraction
time (s)
TKEO
GLR
GLR
GLR
Generalized Likelihood Ratio (GLR):
: actived/non-actived muscle hypothesis.
time (s)
mV
mV
mV
mV
stylohyoid
Posterior digastric
Submandibular
Swalloing sound
A. Mialland et al. Stylohyoid and posterior digastric measurement with intramuscular EMG, submental EMG and swallowing sound. In: BIOSIGNALS (2023)
beginning
Detection limit
Notes: plausibilité

## Slide 27: 3. Statistical study– Data extraction
time (s)
TKEO
GLR
GLR
GLR
Teager-Kaiser Energy Operator (TKEO):
Proportional to the amplitude and to the
frequency content of the signal
time (s)
mV
mV
mV
mV
stylohyoid
Posterior digastric
Submandibular
Swalloing sound
beginning
Detection limit

## Slide 28: 3. Statistical study– Data extraction
temps (s)
mVRMS
RMS amplitude: posterior digastric
time (s)
TKEO
GLR
GLR
GLR
time (s)
mV
mV
mV
mV
stylohyoid
Posterior digastric
Submandibular
Swalloing sound
RMS value of the samples contained
in a sliding window
beginning
Detection limit

## Slide 29: 3. Statistical study– Data extraction
temps (s)
mVRMS
RMS amplitude: posterior digastric
time (s)
TKEO
GLR
GLR
GLR
time (s)
mV
mV
mV
mV
stylohyoid
Posterior digastric
Submandibular
Swalloing sound
)
beginning
Detection limit

## Slide 30: 3. Statistical study– Data extraction
Temporal variables:
Amplitude variables:
temps (s)
mVRMS
RMS amplitude: posterior digastric
time (s)
TKEO
GLR
GLR
GLR
time (s)
mV
mV
mV
mV
stylohyoid
Posterior digastric
Submandibular
Swalloing sound
beginning
Detection limit

## Slide 31: 3. Statistical study– Timing analysis
Time (s)
Start of Swallowing ?
Available
time ?
Tends to be dedicated to swallowing ?
Repeated mesure ANOVA:
  Effect of bolus type ?
  Effect of muscle type ?
A. Mialland et al. Stylohyoid and posterior digastric timing evaluation. In: BSN (2023)

## Slide 32: 3. Statistical study– Timing analysis
Stylohyoid
Posterior Digastric
Submandibular
Temporal limit
means, : p<0,05, : p<0,01
*
**
*
**
*
peaks
Type of bolus
time (s)
saliva
water
thick
solid
Stylohyoid
Posterior Digastric
Submandibular

## Slide 33: 3. Statistical study– Timing analysis
means, : p<0,05, : p<0,01
*
**
*
**
*
peaks
time (s)
onsets
-0,05s
-0,1s
Stylohyoid
Posterior Digastric
Submandibular
Type of bolus
saliva
water
thick
solid

## Slide 34: 3. Statistical study– Recruitment pattern analysis
Time (s)
Start of
Swallowing ?
Available
Time ?
Tends to be dedicated
To swallowing ?
Linear Mixed Models
  Effect of muscle type ?
  Effect of type of task ?
A. Mialland et al. Stylohyoid and posterior digastric recruitment pattern evaluation in swallowing and non-swallowing tasks. In: IRBM (2023). Submitted

## Slide 35: 3. Statistical study– Recruitment pattern analysis
SH
Normalized RMS Amplitude (%)
saliva
water
thick
solid
maximum
Normalized RMS Amplitude (%)
mean
saliva
water
thick
solid
The posterior digastric has the smallest amplitude.
The stylohyoid has no significant variation depending on the type of bolus.
mean
Maximal swallowing
PD
Normalization: for each subject, the RMS signals are divided by the maximum value from all
the swallowing extracted from both the stylohyoid and posterior digastric muscles.

## Slide 36: 3. Statistical study– Recruitment pattern analysis
Acquisitions
Swallowings
mean
Median noise level
Noise: 75th percentile
Normalisation: for each subject, the RMS signals are divided by the maximum value from all
the swallowing, extracted from the currently considered muscle.
Net predisposition for swallowing of the stylohyoid. Trend for the posterior digastric.
No clear tendency for the submandibular muscles,
Amplitude RMS normalisée (%)
stylohyoïdien
Amplitude RMS normalisée (%)
digastrique postérieur
Amplitude RMS normalisée (%)
sous-mandibulaires

## Slide 37: Presentation Plan
Context
Clinical Research Protocol (CRP)
Statistical Study
Real time detection
Implantable Active Artificiel larynx (IAAL)
Conclusion and perspectives

## Slide 38: 4. Real Time Detection – Literature Comparison
Reproduction of similar studies:
Constantinescu et al. 2018:
  Surface EMG: submandibular muscles.
  Swallowing, lips purseing, head movements, tongue movements.
  swallowing, lips purseing, lateral head movements, head flexion/extension.
McNulty et al. 2021:
  Healthy subjects and total laryngectomies
  Surface EMG: submandibular muscles, diaphragm, intercostal muscles.
  Swallowing, coughing, speaking, sit up, sit down, touch head, turn torso, walking.
  Swallowing, coughing, speaking, saying "iii", head movements, head flexion/extension.
Boruta
LDA / SVM
200ms window, step of 50ms
A. Phinyomark et al. 2017.
A. Mialland et al. Stylohyoid and posterior digastric potential evaluation for a real-time swallowing detection, with intramuscular EMG. In: IEEE trans MRB (2023). Submitted

## Slide 39: 4. Real Time Detection – Literature Comparison
Comparing with Constantinescu et al.
precision : 83.9%, recall : 92.3%, F1-score : 87.9%
Comparing with McNulty et al. precision : 77%, recal : 57%, F1-score : 66%
The closest results to the literature.
Substantial increase with the stylohyoid muscle alone.
Potentiel synergestic effect between SH and PD muscles.

## Slide 40: 4. Real Time Detection – Literature Comparison
non-swallowing
swallowing
Method:
Sequential forward selection
Boruta
LDA / SVM

## Slide 41: 4. Real Time Detection – Literature Comparison
Classification performances (F1-score)
Number of tasks added to swallowing
Incrementally adding the best scoring task to classification
The higher the rank, the harder the classification of the taks
LDA: , SVM:
SH: stylohyoid , PD: posterior digastric , SM: submandibular

## Slide 42: 4. Real Time Detection – Literature Comparison
Mastication is always the 13th, causing the biggest drop
Classification performances (F1-score)
LDA: , SVM:
SH: stylohyoid , PD: posterior digastric , SM: submandibular
Number of tasks added to swallowing

## Slide 43: 4. Real Time Detection – Literature Comparison
6
3
2
2
1
rotation de la tête
2,5
Classification performances (F1-score)
LDA: , SVM:
SH: stylohyoid , PD: posterior digastric , SM: submandibular
Number of tasks added to swallowing

## Slide 44: 4. Real Time Detection – Literature Comparison
2,5: lateral head movements
3,2: head flexion/extension
3,7: coughing
5,4: lips purseing
5,5: whistle
5,8: clench teeth
6,6: smilling
7,3: saying « iii »
7,6: blow
9,3: speaking
10,3: jaw opening
10,8: lateral jaw movements
13: mastication
Rang moyen de chaque tâche:
Inclusion rank of the task, for each
classifier and signal combination
From easiest to hardest to classify

## Slide 45: 4. Real Time Detection – Participants subgroups
P1
P2
P3
P4
P5
C4
C3
C1
Boruta
LDA / SVM
Method:
Agglomerative hierarchical clustering
Mean Euclidean distance between swallowing.
P3
C2
C1

## Slide 46: 4. Real Time Detection – Participants subgroups
Effect of subgroups of participants
Greatest increase from the group containing all the participants

## Slide 47: 4. Real Time Detection – Participants subgroups
Normalized % of detected swallowing
Temporal detection limit
All participants
time (s)
Cumulative Frequency
Groupe 1
time (s)
Cumulative Frequency
Groupe 2
time (s)
Cumulative Frequency
Missed muscle
capacity

## Slide 48: Presentation Plan
Context
Clinical Research Protocol (CRP)
Statistical Study
Real time detection
Implantable Active Artificiel larynx (IAAL)
Conclusion and perspectives

## Slide 49: 5. IAAL - Feasibility
Airway reconstruction
Active protective mechanism
[H. C. Grillo 2002] synthetic prosthesis are not effective.
Notes: Stop l’activité biologique des cellules

## Slide 50: 5. IAAL - Feasibility
vue de
A et B
Debry C. et al. 2014.
Airway reconstruction
Notes: Stop l’activité biologique des cellules

## Slide 51: 5. IAAL - Feasibility
Functional tracheal prosthesis [H. Etienne et al. 2018, H. C. Grillo 2002 ]:
Longitudinal flexibility.
Rigidity.
Impermeability: respiratory epithelium.
Non-immunogenicity.
reliability et reproducibility of the method.
Biocompatibility of the materials.
Airway reconstruction
E. Martinod et al. 2018
Notes: Stop l’activité biologique des cellules

## Slide 52: 5. IAAL - Feasibility
Functional protective mechanism:
Increased height of the trachea
Diversion of the bolus
Allow surgical intervention
Compatibility with radiotherapy
Allow expectoration of secretions
Prohibit any permanent closure
Closing speed
Small size
Upper esophageal sphincter
Active protective mechanism

## Slide 53: 5. IAAL - Feasibility
Implantable measurement:
K. A. Yildiz et al. 2020 (survey)
Vue et al. 2020, K. L. Burke et al. 2022:
Active protective mechanism

## Slide 54: 5. IAAL – Laboratory Prototype
Serial communication.
Real-time.
PC
UART
DMA
CPU
Protective
mechanism
Power
Circuit

## Slide 55: 5. IAAL – Laboratory Prototype
PC
UART
DMA
CPU
Protective
mechanism
Power
Circuit

## Slide 56: 5. IAAL – Laboratory Prototype
PC
UART
DMA
CPU
Protective
mechanism
Power
Circuit
15mm
20mm
10mm
39mm
65mm

## Slide 57: 5. IAAL – Laboratory Prototype
Evaluated Configurations:
Classifiers: LDA, SVM, ANN
Window size: 200 samples, 400 samples
Window step: 50 ms
Signal combinations: SH-PD, SH-PD-SM
Microcontrollers tested

## Slide 58: 5. IAAL – Laboratory Prototype
Results per configuration
Speed camera (1000 fps):
Swallowing detection turns a LED on.
Total time worst case - 400 samples, SH-PD-SM, ANN:
+ 3,27 = 25,27ms

## Slide 59: Presentation Plan
Context
Clinical Research Protocol (CRP)
Statistical Study
Real time detection
Implantable Active Artificiel larynx (IAAL)
Conclusion and perspectives

## Slide 60: 6. Conclusion and perspectives
Stylohyoid (SH) and posterior digastric (PD) muscles measurement:
  Innovative measurement method with intramuscular electromyography.
Timings:
  SH and PD muscles activates at the beginning of swallowing.
  Earliness of PD muscle activity.
Amplitude:
  Predisposition of SH muscle to swallowing.
  Stability of SH muscle.

## Slide 61: 6. Conclusion and perspectives
Real time detection:
  Significant improvement in detection Accuracy with SH muscle alone.
  Significant improvement in Earliness of detection with PD muscle alone.
  Potential synergistic effect between SH and PD muscles.
Implantable active artificial larynx:
  Measurement and airway reconstruction methods are promising.
  Laboratory prototype actuated in 25,27ms.
Beginning of swallowing
0 secondes
Limite de détection
Interval of activation of
the measured muscles
Signal processing
()
Actuation of the
System ()
Moment of activation of
Measured muscles
temps

## Slide 62: 6. Conclusion and perspectives
Research of additional information on neck activity.
Development of a active protective mechanism of the trachea.
Evaluated the hypotheses of total laryngectomees.
Elaboration of a algorithm with strong temporal constraints.
Evaluation of corner cases: vomiting, sleep, evolution through time, …

## Slide 63: Publications
International peer-reviewed journals:
  Survey
  A. Mialland, I. Atallah et A. Bonvilain. “Toward a robust swallowing detection for an implantable active artificial larynx : a survey”. In : Medical & Biological Engineering & Computing (2023), p. 1299-1327. doi : 10.1007/s11517-023-02772-8.
  Articles
  A. Mialland, I. Atallah et A. Bonvilain. “Stylohyoid and posterior digastric potential evaluation for a real-time swallowing detection, with intramuscular EMG”. In: IEEE Transactions on Medical Robotics and Bionics (2023). Doi: 10.1109/TMRB.2023.3336960.
  A. Mialland, I. Atallah et A. Bonvilain. “Stylohyoid and posterior digastric recruitment pattern evaluation in swallowing and non-swallowing tasks”. In : Innovation and Research in BioMedical engineering (2023). Doi: 10.1016/j.irbm.2024.100823.
  A. Mialland et al. “Submental MechanoMyoGraphy (MMG) to Characterize the Swallowing Signature”. In : Innovation and Research in BioMedical engineering (2021). doi : 10.1016/j.irbm.2021.05.001.
International peer-reviewed conference:
  A. Mialland, I. Atallah et A. Bonvilain. “Stylohyoid and posterior digastric measurement with intramuscular EMG, submental EMG and swallowing sound.” In :Biomedical Engineering Systems and Technologies (2023). Lisbonne. doi : 10.5220/0011628100003414.
  A. Mialland, I. Atallah et A. Bonvilain. “Stylohyoid and posterior digastric timing evaluation”. In : Body Sensor Networks (2023). Boston. Doi: 10.1109/BSN58485.2023.10331308.
Internationale conference:
  A. Mialland, I. Atallah et A. Bonvilain. The inherent complexity of an implantable active artificial larynx. Recherche en Imagerie et Technologies pour la Santé. Brest. 2022.
Nationale conference:
  A. Mialland, I. Atallah et A. Bonvilain. Vers la faisabilité d’un larynx artificiel implantable actif. Société Française de Phoniatrie et Laryngologie. Tours. 2021.
Clinical research Protocol:
  A. Mialland, I. Atallah et A. Bonvilain. Detection of Early Swallowing Time by Electromyogram and Sound Recording in Healthy Volunteers. https://clinicaltrials.gov/. Protocol ID : 38RC22.0096.2022.